SOUP DOCUMENTATION — QR Webcam Scanner

Company: Precision Biometrics, Inc.
Product: MyoVision Software

1. Identification

SOUP Name:

ZXing-C++ (barcode decoding library)

OpenCV (computer vision library)

Version Used:

- ZXing-C++: 2.x (validated build)

- OpenCV: 4.x (validated build)

Usage Location:

- PBIQrWebcamScanner.exe (external helper application)

- Not embedded directly in core VC6 application

2. Intended Use

The SOUP components are used to:

Capture webcam images (OpenCV)

Decode QR/DataMatrix barcodes (ZXing)

Extract device serial number (custom logic)

Output:

Serial number written to:

C:\Temp\PBI_QR_Result.ini

3. Safety Classification

Function: Non-diagnostic support

Impact: Incorrect device identification

Risk Level: Low–Moderate

4. Risk Analysis

Hazard: Incorrect serial number captured

Causes:

Poor lighting

Damaged label

Decode error

Partial read

Mitigations:

Validate serial format (length, pattern)

Display serial to user for confirmation

Allow manual entry override

Timeout / retry if decode fails

Reject empty or malformed results

5. SOUP Evaluation

Widely used: Yes

Active maintenance: Yes

Known defects reviewed: Yes

License acceptable: Yes (Apache 2.0 / BSD)

6. Verification Testing

Test Types:

Known good labels (multiple devices)

Low light conditions

Glare conditions

Invalid QR codes

No QR present

Acceptance Criteria:

Correct serial extracted ≥ 99% of valid scans

Invalid scans rejected (no silent success)

No crash / hang conditions

7. Configuration Management

Versions are fixed and documented

Libraries bundled with release build

No runtime updates

8. Update Policy

Updates require:

Regression testing

Re-validation of scan accuracy

Documentation update

9. Failure Handling

Camera failure → Error returned

Decode failure → Retry / timeout

Invalid serial → Reject + user prompt

Timeout → User retry option

10. Cybersecurity Considerations

No network access

Local-only processing

No PHI/PII handled

11. Traceability

Capture serial → ZXing decode

Validate serial → Custom parser

User confirmation → UI display

Manual fallback → Text input